11D-8.002 Definitions.

(1) Acceptable Range – the results of alcohol reference solutions and dry gas standard analyses which fall within the following ranges at each alcohol vapor concentration: 0.05 g/210L range is 0.045 to 0.055 g/210L; 0.08 g/210L range is 0.075 to 0.085 g/210L; 0.20 g/210L range is 0.190 to 0.210 g/210L; or the Alcohol Reference Solution gas chromatographic results which fall within the following ranges: 0.0605 g/100mL range is 0.0586 to 0.0623 g/100mL; 0.0968 g/100 mL range is 0.0938 to 0.0997 g/100mL; 0.2420 g/100mL range is 0.2347 to 0.2492 g/100mL.

(2) Accuracy – the nearness of a measurement to a known concentration.

(3) Acetone Stock Solution – a mixture of acetone and distilled or deionized water provided by the Department.

(4) Agency – a law enforcement agency other than the Department, or an entity which conducts breath tests or submits blood samples for alcohol testing pursuant to these rules, or a civilian entity performing such duties on behalf of a law enforcement agency.

(5) Agency Inspection – the periodic testing of the calibration and operation of a breath test instrument, including all required preventive maintenance, in accordance with Rule 11D-8.006, F.A.C., and performed by a person authorized by the Department.

(6) Agency Inspector – a person who has been issued an Agency Inspector permit by the Department.

(7) Alcohol – ethyl alcohol, also known as ethanol.

(8) Alcohol Free Test – a result of 0.000 g/210L when using distilled or deionized water.

(9) Alcohol Reference Solution – a standard used to verify the calibration of a breath test instrument consisting of a mixture of alcohol and distilled or deionized water that will produce a known alcohol vapor concentration at a specific temperature.

(10) Analyst – a person who has been issued a permit by the Department to conduct blood alcohol analyses.

(11) Approved Blood Alcohol Test – the analyses of two separate portions of the same blood sample using a Department-approved blood alcohol test method and a Department-approved procedure, with results within 0.010 grams of alcohol per 100 milliliters of blood (g/100mL), and reported as the blood alcohol level.

(12) Approved Breath Alcohol Test – a minimum of two samples of breath collected within fifteen minutes of each other, analyzed using an approved breath test instrument, producing two results within 0.020 g/210L, and reported as the breath alcohol level. If the results of the first and second samples are more than 0.020 g/210L apart, a third sample shall be analyzed. Refusal or failure to provide the required number of valid breath samples constitutes a refusal to submit to the breath test. Notwithstanding the foregoing sentence, the result(s) obtained, if proved to be reliable, shall be acceptable as a valid breath alcohol level.

(13) Authorized Repair Facility – the Department, the breath test instrument manufacturer, an entity authorized by the breath test instrument manufacturer to service and repair such breath test instrument.

(14) Blood – whole blood.

(15) Blood Alcohol Level – the alcohol concentration by weight in a person’s blood based upon grams of alcohol per 100 milliliters of blood (g/100mL).

(16) Breath Alcohol Level – the alcohol concentration by weight in a person’s breath based upon grams of alcohol per 210 liters of breath (g/210L).

(17) Breath Test Instructor – a person who has been issued a Breath Test Instructor Certification by the Criminal Justice Standards and Training Commission.

(18) Breath Test Operator – a person who has been issued a Breath Test Operator permit by the Department.

(19) Department – the Florida Department of Law Enforcement.

(20) Dry Gas Standard – a standard consisting of a mixture of alcohol and gas which produces a known alcohol vapor concentration used to verify the calibration of a breath test instrument.

(21) Evidentiary Breath Test Instrument – a breath test instrument approved by the Department under Rule 11D-8.003, F.A.C., and used primarily to conduct alcohol breath tests pursuant to Florida law.

(22) Methods – types of alcohol analyses approved by the Department to conduct chemical or physical tests of blood or breath.

(23) Mouth Alcohol Solution – a mixture of alcohol and distilled or deionized water provided by the Department.

(24) Permit – when issued by the Department, certifies that the holder has met all necessary qualifications, remains in full compliance with these rules and is authorized to perform all related duties. A permit is issued only to a qualified applicant and remains valid and in full effect until determined otherwise by the Department.

(25) Reference Sample Device – a device, also known as a simulator, that produces a known vapor concentration by the passage of air through a liquid.

(26) Target Concentration – a gas chromatographic result equivalent to the following known alcohol vapor concentrations of alcohol reference solution: for 0.05 g/210L the target concentration is 0.0605 g/100mL; for 0.08 g/210L the target concentration is 0.0968 g/100mL; for 0.20 g/210L the target concentration is 0.2420 g/100mL.

Specific Authority 316.1932(1)(a)2., (f)1., 316.1933(2)(b), 316.1934(3), 322.63(3)(a), (b), 327.352(1)(b)3. FS. Law Implemented 316.1932(1)(b)2., 316.1933(2)(b), 316.1934(3), 322.63(3)(b), 327.352(1)(e), 327.353(2), 327.354(3) FS. History–New 10-31-93, Amended 1-1-97, 7-6-99, 7-29-01, 11-5-02, 12-9-04, 3-27-06.